| RGA300 amu RGA (with RS-232 cable & free software) |  |
| --- | --- |
| Mass range RGA300 | 1 to 300 amu |
| Mass filter type | Quadrupole |
| Detector type | Faraday cup — std. electron multiplier — opt. |
| Resolution | Better than 0.5 amu @ 10 % peak height (per AVS std. 2.3). Adjustable to constant peak width throughout the mass range. |
| Sensitivity (A/Torr) | 2 × 10-4 (FC),<200 (EM). User adjustable throughout high voltage range. Measured with N2 @ 28 amu with 1 amu full peak width, 10 % height, 70 eV electron energy, 12 eV ion energy, and 1 mA electron emission current. |
| Min. detectable partial pressure | 5 × 10-11 Torr (FC). 5 × 10-14 Torr (EM). Measured with N2 @ 28 amu with 1 amu full peak width, 10 % height, 70 eV electron energy, 12 eV ion energy, and 1 mA electron emission current. |
| Operating pressure | 10-4 Torr to UHV (FC) 10-6 Torr to UHV (EM) |
| Max. operating temp. | 70 °C |
| Bakeout temperature | 300 °C (without ECU) |
| Ionizer |  |
| Design | Open ion source, cylindrical symmetry, electron impact ionization. |
| Material | SS304 construction |
| Filament | Thoriated-iridium (dual) with firmware protection. Built-in 1 to 10 W degas ramp-up. Field replaceable. |
| Electron energy | 25 to 105 V, programmable |
| Ion energy | 8 or 12 V, programmable |
| Focus voltage | 0 to 150 V, programmable |
| Electron emission current | 0 to 3.5 mA, programmable |
| General |  |
| Probe dimension | 8.75" from flange face to top of ionizer |
| Probe insertion | 2.0" |
| Probe mounting flange | 2.75" CF |
| Minimum tube I.D. | 1.375" |
| ECU dimensions | 9.1" × 4.1" × 3.1" (WHL) Easily separated from the probe for bakeout. |
| LED indicators | Power ON/OFF, filament ON/OFF, degas ON/OFF, electron multiplier ON/OFF, RS-232 Busy, Error, Overpressure, Burnt Filament. |
| Warm-up time | Mass stability ±0.1 amu after 30 minutes |
| Computer interface | RS-232C, 28,800 baud with high-level command set |
| Software | Windows based application |
| Power requirement | 24 VDC @ 2.5 amps. Optional built-in VAC adapter. |
| Weight | 6 lbs. |
| Warranty | One year parts and labor on defects in materials and workmanship |